FORCE PROVING INSTRUMENT

{{CASE_FILE}}/D1.05/{{CERTIFICATE_NO}}

| {{CALIBRATION_DATE}} |  | {{EXPIRY_DATE}} |  | 1 |  | 2 |
| --- | --- | --- | --- | --- | --- | --- |

| 1. Calibrated for: | {{CUSTOMER_NAME}} |
| --- | --- |
|  | {{CUSTOMER_ADDRESS}} |

Customer Ref. No.: {{CUSTOMER_REF_NO}} Dated: {{CUSTOMER_REF_DATE}}

| 2. Description & identification | Type: {{RING_TYPE}} | Digital Indicator Sr. No: {{DIGITAL_INDICATOR_SRNO}} |
| --- | --- | --- |
| of instrument | Capacity:{{CAPACITY}} | Manufacturer: {{MANUFACTURER}} |
|  | Sl. No.: {{EQUIPMENT_SRNO}} | Resolution: {{RESOLUTION}} {{RESOLUTION_UNIT}} |
|  | Connector Type:{{CONNECTOR_TYPE}} pin | Cable Length: {{CABLE_LENGTH}} m |
|  | Model: {{MODEL_NO}} | Make: {{MAKE}} |

Accessories:  Self -aligning {{TYPE_FORCE}}

| ##### Temperature: | ##### (23 1) C | ##### Relative humidity: (50 10)% |
| --- | --- | --- |

##### 3. Environmental conditions:

##### 4. Standard(s) used with associated:	{{FORCE_MACHINE}}

##### measurement uncertainty 		±{{BMC}}% (k=2)

##### 5. Metrological traceability of standard(s) used: 	The standard(s) used for calibration is(are) traceable to the National Standard, which realize the units of quantities according to the International System of Units (SI).

# 6. Principle/methodology of calibration:       Sub-Div.#D1.05/Doc.#3/CP#FT/F-02

# and calibration procedure number:         broadly based on IS 4169-2014 (ISO:376-2011).

# No load output: The digital indicator was switched on for 30 minutes to warm up and stabilized for no

# load output before the start of calibration.  The no load output was noted (before taring) and the calibration

# signal was noted

Preloading:	Before the application of the calibration forces, the instrument was preloaded thrice to its maximum capacity and kept at full load for about 90 seconds.

Calibration:	The sequence of the applied calibration force in {{TYPE_FORCE1}} is given below:

At 0: Two series of calibration forces in increasing values.  At {{POSITION_TEXT}} positions: One series of calibration forces each in increasing values. Creep test is performed by calculating the difference in output i30 obtained at 30s and i300 obtained at 300s after the removal of the maximum calibration force and express this difference as percentage of maximum deflection.

The calibration was made by using self-aligning {{TYPE_FORCE}} provided along with the instruments to ensure axial application of the force.

Between each series, the instrument was rotated along its axis so as to occupy during the calibration three positions  (0, {{POSITION_TEXT}}) and the instrument was subjected to the full load once for about 90 seconds each time before starting in a new position. Between the loadings, readings corresponding to no load after waiting at least 30 seconds for the return to zero were noted. Prior to each reading the dial gauge was lightly tapped at the center of the dial cover.

FORCE PROVING INSTRUMENT

{{CASE_FILE}}/D1.05/{{CERTIFICATE_NO}}

| {{CALIBRATION_DATE}} |  | {{EXPIRY_DATE}} |  | 2 |  | 2 |
| --- | --- | --- | --- | --- | --- | --- |

7. Results: {{TYPE_FORCE1}}

| ##### The calibration data obtained in compression is valid for the following digital indicator setting only |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| ##### Calibration Signal: Nil |  |  | ##### No Load Output: {{NO_LOAD_OUTPUT}} {{NO_LOAD_OUTPUT_UNIT}} |  |  |
|  | ##### (Digital Indicator Reading in {{NO_LOAD_OUTPUT_UNIT}}) |  |  |  |  |
| ##### Applied ##### Force kN | ##### Position 0 ##### series 1 | ##### Position 0 ##### series 2 | ##### {{POS3}} series 3 | ##### {{POS4}} series 4 | ##### Average 1,3,4 |
| ##### {{DATA_ROW}} |  |  |  |  |  |

Interpolation Equation:

F = {{AVG_TO_FORCE_EQ}}

X = {{FORCE_TO_AVG_EQ}}

Where F = Force in kN

X = Indicator reading in {{RESOLUTION_UNIT}}

##### Classification:    The force proving instrument is found to comply with the requirements of IS: 4169-2014				(ISO: 376-2011) in respect of specific forces and classified as follows:

| ##### Class | ##### Mode | ##### From | ##### To | ##### Uncertainty of Measurement |
| --- | --- | --- | --- | --- |
| {{CLASSIFICATION_TABLE}} |  |  |  |  |

The reported uncertainty is at coverage factor k=2 which corresponds to a coverage probability of approximately 95% for a normal distribution, considering the relative deviation of different components such as zero, repeatability, reproducibility, resolution, creep and combining with the uncertainty of the applied force.

8. Date(s) of calibration:   {{CALIBRATION_DATE}}

9. Remarks:  {{REMARKS}}